CEM Microwave Manual

- Place all reagents into 10 mL, or 35 mL vial with a stir bar.

- Seal the vessel using the supplied cap, with the notches positioned downward.

- Position cap above vessel with the notches pointing down.

- Place cap on top of vessel at a slight angle and roll your finger over the top of the vessel.

- Press down until a click is audible, indicating the top is properly positioned.

- Ensure that cavity is clean and empty.

- Open the Synergy software.

- Parameters:

  - Mode – Explorer

  - Explorer – Explorer 12

  - Explorer – Hybrid rack : 2(35 ml) and 5(10 ml)

  - Options – Temperature – Infrared

- With the rack locked into position in the autosampler tray, place the vial(s) into the Explorer rack.

- Select (highlight) the method to be added to the rack.

- Details regarding the method types and method programming are outlined in the Explorer Operation Manual.

- “Double click” the desired rack position in which to add the method.

- Press the PLAY key located at the top of the Synergy screen.

- The autosampler will automatically load the selected vessel(s) and send the method parameters to Discover.

- The pressure device will automatically move into position prior the microwave radiation.

- Method parameters can be changed during a reaction by using the hot keys.

Removing the vessel:

- When the reaction is complete, the reaction vessel will cool until the release parameters are achieved. Once these parameters are attained, the pressure device automatically releases and moves to its home position. The sample is returned to the Explorer rack.

- If the APD does not automatically release the reaction vessel, the pressure (or temperature) is above the release limit. A message will appear indicating the current temperature and pressure value. Cool the reaction vessel completely, and manually release the APD. Once the APD is in its home position, remove the vessel from microwave cavity, and place it in a fume hood or other ventilated area. Caution: To prevent the possibility of severe burns, wear insulate gloves and protective gear as outlined in the user’s safety program. Contents of the vail(s) may be under pressure and should be released within a fume hood.

- After the vial has been returned to the Explorer rack, the cap can be removed and work up can be performed.